YEAR-2017

- Urgent Field Safety Notice for MX40 IntelliVue by Philips Medical Systems

Date - 2017.08.07
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

Philips has recently discovered that under specific 802.11 network conditions, a partial reassociation of the MX40 WLAN monitor to a compatible central monitoring system (“Information Center”) is possible. In this state, although the Information Center provides a visible and audible “No Data Tele” INOP alert, the MX40 WLAN itself enters telemetry mode, i.e., its screen turns off in one minute and local alarming is disabled.

HOW TO IDENTIFY AFFECTED PRODUCTS:

- The Product Number and exchange part number are contained on the device’s product label, located on the back of the device.

- The product number is indicated after the symbol and may be either 6- numbers or 12-numbers long.

- The exchange part number is indicated after the Service #.

- The MX40 software revision is identified by touching the battery icon in the lower right of the MX40 display which opens the Device Status window.

- In the Device Status window, touch the Device Info text, to display the application software (Appl SW) version.

ACTIONS PLANNED BY PHILIPS:

- Philips is releasing a software upgrade B.06.18 for affected Philips IntelliVue MX40 Patient Wearable Monitors.

- A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade.

- Urgent Field Safety Notice for Philips IntelliVue NMT Patient Cables 98980317458 by Philips Medical Systems

Date - 2017.05.22
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

A small number of IntelliVue NMT Patient Cables may contain a localized isolation defect in the shielding of the acceleration sensor housing. This manufacturing defect may cause localized heating due to an unintended current flow between the acceleration sensor housing attached to the patient thumb and the NMT stimulation electrode connection attached to the wrist of the same hand.

HOW TO IDENTIFY AFFECTED PRODUCTS:

- Check the Notified Body number which amends the CE mark on the NMT Patient Cable.

- All affected NMT Patient Cables bear the CE Mark with the suffix 0366: 0366.

- Currently shipped NMT Patient Cables with CE mark and suffix 0123 are not affected.

- These have already been modified and do not need to be replaced.

ACTIONS PLANNED BY PHILIPS:

Philips is voluntary initiating a correction consisting of:

• Distribution of this Field Safety Notice (FSN).

• Replacement of affected NMT Patient cables.

A Philips Healthcare representative or authorized service provider will contact customers to arrange for a visit to exchange all affected cables at the customer site.

YEAR-2018

- Urgent Field Safety Notice for MX40 IntelliVue by Philips Medical Systems

Date - 2018.05.28
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

Five warning statements are missing from the IntelliVue MX40 IFU for software revisions B.05, B.06 and B.06.5X.

HOW TO IDENTIFY AFFECTED PRODUCTS:

IntelliVue MX40 Instructions for Use for software revisions B.05, B.06 and B.06.5X.

ACTIONS PLANNED BY PHILIPS:

Philips is voluntarily initiating a correction consisting of:

- Distribution of this Field Safety Notice (FSN86201829A).

- Release of an IFU Errata that provides the missing warnings A Philips Healthcare            representative will contact customers with affected IFU revisions to provide a copy of the errata.

This FSN includes the Errata providing the missing warning statements.

- Urgent Field Safety Notice for Philips IntelliVue Information Center iX by Philips Medical Systems

Date - 2018.01.23
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

The issue affects all PIIC iX Surveillance and Patient Link revisions. The issue will occur after a surveillance station is first restarted during the year 2018. After that initial restart, the station will thereafter no longer perform patient Discharge and Transfer operations.

If any of the following Discharge or Transfer operations are attempted, the operation will not be completed and, instead, the Surveillance Station will restart:

• Discharging a patient under certain conditions

• Transferring a patient (from bedside or central station)

• Resolving patient conflicts (from bedside or central station) where bedside’s patient must overwrite PIIC iX’s patient (revision B.0x and C.0x only)

• Clearing a surveillance sector

• Assigning a bed to an empty surveillance sector The Surveillance station will continue to monitor patients after any restart.

HOW TO IDENTIFY AFFECTED PRODUCTS:

PIIC iX revision A.0X or B.0X or C.0: Identify the affected product by clicking on the Philips icon on the PIIC iX surveillance display.

This will bring up the Product Support Page which identifies the product.

ACTIONS PLANNED BY PHILIPS:

- Philips has initiated a correction to address this issue and will provide this software update to customers with affected devices at no charge.

- Software updates for PIIC iX A.0x and B.0x are expected to be available by January 8 th .

- A Philips Healthcare representative will contact customers with affected devices to arrange for correction of the issue.

YEAR-2019

- URGENT - Medical Device Recall IntelliVue MX4O

Date - 2019.05.10
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

The MX4O may experience increased power consumption and may have a lack of visual and/or audible notifications while in a “Low Battery State”.

HOW TO IDENTIFY ACTION TO BE TAKEN:

This action applies to all MX4O’s. AFFECTED PRODUCTS

ACTIONS PLANNED BY:

- Philips Healthcare will release a software upgrade, which will correct this PHILIPS problem. Philips will contact you to arrange for software correction.

YEAR-2020

- Urgent Field Safety Notice for IntelliVue Multi measurement Server X2 by Philips Medizin Systeme Boblingen GmbH

Date - 2020.12.16
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

Speakers on lntelliVue X2 and lntelliVue MP2 may fail. This failure may be identified by the display of a "Speaker Malfunct." INOP or by the absence of audible sound.

HOW TO IDENTIFY AFFECTED PRODUCTS:

The serial number of the Philips M3002A (865039) lntelliVue X2 and M81 02A (865040) lntelliVue MP2 is contained on the product Iabei, located on the rear of the device.

ACTIONS PLANNED BY PHILIPS:

- Philips will replace affected speaker assemblies free of charge.

- PHIUPS A Philips Healthcare representative will contact you to arrange for the replacement of the speaker assembly in the listed affected devices.

- Urgent Field Safety Notice for TRx4851A 2.4 GHz / Trx4841A 1.4 GHz IntelliVue Tel. by Philips Medical Systems

Date - 2020.03.23
Product group - Medical electronics / Electromedical devices (Electrodiagnostics)

PROBLEM DESCRIPTION:

The ECG signal from patients being monitored using a Philips TRx4841 A and TRx4851A Telemetry Transceiver may not be properly analyzed when it is used with a; Philips Patient Information Center iX Release C.02.xx or C.03.01. lf this occurs, the Information Center will not display a heart rate or generate, display, or annunciate any heart rate or arrhythmia alarms. Alarming based on pulse oximetry (Sp02) signals are unaffected by this issue.

HOW TO IDENTIFY AFFECTED PRODUCTS:

- The Philips IntelliVue TRx4841A and TRx4851A Telemetry Transceivers can be identified by the label on the front of the device:

- To identify the revision of software on the Philips Patient Information Center iX surveillance: Access the Product Support screen from the Main Setup task bar button or by clicking the Philips icon.

- Philips TRx4841A and TRx4851A Telemetry Transceivers, when used with Patient Information Center iX C.02.xx or C.03.01 are at risk of potentially exhibiting this behavior.

- Philips has confirmed that this issue is not present when Philips TRx4841A and TRx4851A Telemetry Transceivers are used with PIIC Glassie N.01 .22, PIIC iX A.02.16, or PIIC iX B.02.18.

ACTIONS PLANNED BY PHILIPS:

Philips will contact you should further action be necessary.