## Sheet1
Manufacturer | Brand Name | Device Problem | Patient Problem | Event date | Event type | Manufacturer Narrative | Event Description
Philips Medical Systems | IntelliVue MX700 Patient Monitor | Therapeutic or Diagnostic Output Failure (3023) | Death (1802) | 2020-12-12 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that the intellivue mx700 patient monitor went into standby on (b)(6) 2020. The death of a patient was reported. No further information regarding the patient, such as gender, age, height, and weight, had been made available at the time the reporting decision was due.
Philips Medical Systems | IntelliVue MX800 Patient Monitor | Device Alarm System (1012) | Death (1802) | 2020-12-04 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that the intellivue mx800 patient monitor did not alarm for a desat on (b)(6) 2020 at 10:00 for the patient in icu (b)(6). The patient passed away on (b)(6) 2020.
Philips Medical Systems | IntelliVue MX700 Patient Monitor | Detachment of Device or Device Component (2907) | Death (1802) | 2020-12-05 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that no abp low red alarm was triggered from the intellivue mx700 patient monitor after patient torn off patients cable on (b)(6) 2020, bed id 9b-3e between 19:30 - 21:00. The patient died.
Philips Medical Systems | IntelliVue Information Center IX | Therapeutic or Diagnostic Output Failure (3023) | Cardiopulmonary Arrest (1765); Death (1802) | 2020-12-07 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that on (b)(6) 2020, a patient went into cardiac arrest and expired. The patient was assigned to tele #11 in rm 202 in the med/surg unit. The customer reported having trouble with maintaining the ecg rhythm on a philips information center ix (pic ix) / m4841 telemetry system.
Philips Medical Systems | IntelliVue Information Center Client | Device Alarm System (1012) | Death (1802) | 2020-12-04 | Death | A follow-up report will be submitted once the investigation is complete. | The customer reported an alarm did not go off, and they couldn't print rhythm information for time of even or before the event; the patient passed away.
Philips Medical Systems | IntelliVue Information Center IX | Defective Alarm (1014) | Death (1802) | 2020-11-17 | Death | A follow-up report will be submitted once the investigation is complete. | The customer reported that on (b)(6) 2020 at 0818 a patient in bed (b)(6) had a ventricular fibrillation (vfib) event that did not alarm and the patient expired.
Philips Medical Systems | IntelliVue Information Center IX | Device Alarm System (1012); Communication or Transmission Problem (2896) | Cardiac Arrest (1762); Death (1802) | 2020-12-02 | Death | A follow-up report will be submitted once the investigation is complete. | The customer reported a lack of transmission of bed tele37 to the philips information center ix (pic ix) central station for ten minutes resulting in a missed asystole alarm. The deterioration of the patient's condition and the cardiovascular arrest were not recognized in time, and as a result the a patient died.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Adverse Event Without Identified Device or Use Problem (2993) | Cardiac Arrest (1762); Death (1802) | 2020-11-06 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported a cardiac arrest of a patient with an incidence of death and wanted to know if the monitor sounded correctly. The patient died.
Philips Medical Systems | IntelliVue Information Center IX | Alarm Not Visible (1022) | Death (1802) | 2020-11-13 | Death | This is a supplemental report to mfr. Report #:9610816-2020-00431. the fda registration number initially chosen was incorrect and product listed was incorrect. Please use this report. | The customer reported that one monitor has not alarmed and the patient died
Philips Medical Systems | IntelliVue Information Center IX | Insufficient Information (3190) | Death (1802) | 2020-11-11 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported a patient death on (b)(6) 2020 and wanted to have the logs pulled to determine what happened.
Philips Medical Systems | MP30 IntelliVue Patient Monitor | Device Alarm System (1012) | Death (1802) | 2020-09-27 | Death | Philips field service engineer (fse) went to the customer site and spoke with the customer biomed who complained about the lack of red alarm (asystole) signals from the monitor. The fse performed an alarm simulation test on the monitor and was not able to reproduce the alarm failure. The fse discussed the findings with the customer biomed and determined the setting for the ecg alarm was turned off by a member of the hospital staff. There was no product malfunction. The results of these findings confirmed the device was working as expected. | The customer reported on (b)(6) 2020, their intellivue mp30 monitor failed to alarm. The patient died.
Philips Medical Systems | IntelliVue MX40 802.11A/B/G | Insufficient Information (3190) | Death (1802) | 2020-09-07 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that on (b)(6) 2020 a patient expired while being monitored with the patient worn device. The hospital was investigating the cardiac rhythm of the patient prior to the patient's death
Philips Medical Systems | IntelliVue X3 | Therapeutic or Diagnostic Output Failure (3023) | Death (1802) | 2020-10-19 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that on (b)(6) 2020 the spo2 was not working. The patient died.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Adverse Event Without Identified Device or Use Problem (2993) | Death (1802); Hematoma (1884); Hemorrhage, Subarachnoid (1893); High Blood Pressure/ Hypertension (1908); Irregular Pulse (2469 | 2020-07-21 | Death | A follow up report will be submitted once the investigation is complete. | Via medwatch mw5095852, the customer reported that on (b)(6) 2020, a patient was admitted to the icu due to grade 5 subarachnoid hemorrhage. A ventriculostomy drain was placed and throughout the day, the patient had high/abnormal blood pressures with very narrow pulse pressures. The patient had a subdural hematoma (sdh) and a cardene infusion was being titrated according to the non-invasive blood pressure (nbp) cuff measurements. Around 14:00, the patient went into asystole and passed away shortly after.
Philips Medical Systems | MP90 IntelliVue Patient Monitor | Alarm Not Visible (1022) | Death (1802) | 2020-08-21 | Death | A follow up report will be submitted once the investigation is complete | The customer called philips to report that their mp90 bedside monitor failed to alarm and that the customer requested a log review of the alleged event. A patient death has been reported. The customer has also reported that they feel like there was a delay in treatment because the mp90 did not alarm.
Philips Medical Systems | MP70 IntelliVue Patient Monitor | Alarm Not Visible (1022) | Death (1802) | 2020-08-06 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported a sentinel event that occurred on (b)(6) 2020, and their mp70 intellivue patient monitor failed to alarm. The patient died.
Philips Medical Systems | IntelliVue Info Center IX | No Display/Image (1183) | Death (1802) | 2020-06-25 | Death | A follow up report will be submitted once the investigation is complete. | The customer reported that they were monitoring spo2 only for a patient, but they were unable to see the patient vitals at the philips information center ix (pic ix). The patient was noted to have died.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Device Fell (4014) | Injury (2348) | 2021-01-28 | Injury | A follow up report will be submitted once the investigation is complete. | The customer reported that the intellivue multi measurement server x2 somehow fell down from a mount. The device fell onto a housekeeper, who was injured. No further info was made available at the time the reporting decision was due
Philips Medical Systems | IntelliVue MX800 Patient Monitor | No Audible Alarm (1019) | Cardiopulmonary Arrest (1765); Low Oxygen Saturation (2477) | 2021-01-25 | Injury | The reported issue was investigated via remote support by a philips remote clinical applications support specialist (rcass). According to the biomed at the customer site, a member of the clinical staff was standing in front of the mx800 and saw that the spo2 value was dropping below the desat limit before returning to a level above the desat limit violation. The staff member expected a desat alarm for the drop below the desat limit which was not announced by the monitor. The rcass reviewed the audit log from the central station together with the biomed and multiple yellow spo2 alarms as well as some red desat alarms were found for the patient. The rcass advised the biomed of the alarm delay configuration of 10 seconds for a spo2 low limit alarm and 20 seconds for a desat alarm. The spo2 value returned above the desat limit threshold within the alarm delay time, which was the reason why the alarm was not indicated. The biomed stated that the clinical staff was not aware of the alarm delay functionality and that she would instruct the staff accordingly. The rcass provided the biomed with information regarding the alarm behavior to be forwarded to the clinical users. The biomed was also instructed on how to access and review the alarm review on the monitor. Additionally, the biomed advised that she had tested the monitor and found it to be working as intended. It is considered that this issue was resolved by the remote clinical application support specialist instructing the biomed regarding the alarm delay functionality and providing information regarding the alarm behavior. The monitor worked as intended and there was no malfunction of the device. This issue was caused by the alarm delay configuration. | The customer reported that the intellivue mx800 patient monitor did not alarm for a desat as expected on (b)(6) 2021. The patient needed to be resuscitated.
Philips Medical Systems | IntelliVue Information Center IX | Inadequate Instructions for Healthcare Professional (1319) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-17 | Injury | A philips response center engineer (rce) spoke to the customer biomedical engineer (biomed). The biomed needed to review strips from the patient. The rce instructed the biomed how to readmit patient into pic ix in order to review strips, and advised the biomed to call back if further assistance is needed as the biomed not in front of system at the time. No further response was received from the customer. There was no product malfunction reported; the biomed was looking for assistance obtaining patient strips for review. The type of patient incident was not able to be verified, as no additional information was received from the biomed | The customer reported a patient incident and wanting to obtain information from the philips information center ix (pic ix); they wanted to know how to review data on patient from (b)(6) 2020. Patient incident was reported, no information was available upon request
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | No Display/Image (1183); Therapeutic or Diagnostic Output Failure (3023) | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available (4581) | 2020-11-26 | Injury | A follow-up report will be submitted once the investigation is complete. There is currently insufficient information available to confirm the nature of the alleged serious injury (patient required cpr), and the relationship to the complaint device. | The customer reported that their x2 lost its display for about five minutes and the device could not be restarted during that time. When the x2 was restarted, the nurses determined that the patient needed cardiopulmonary resuscitation (cpr)
Philips Medical Systems | IntelliVue | Break (1069); Device Fell (4014) | Insufficient Information (4580) | 2020-10-29 | Injury | A follow up report will be submitted once the investigation is complete. | The customer reported the clips holding their intellivue mmx device were broken. The mmx device fell and struck a patient in the head.
Philips Medical Systems | IntelliVue MX700 Patient Monitor | Adverse Event Without Identified Device or Use Problem (2993) | Hypoxia (1918) | 2019-12-11 | Injury | A follow up report will be submitted once the investigation is complete. | The customer reported that on (b)(6) 2019 a patient became hypoxic and it is not clear at this time if the alarm sounded
Philips Medical Systems | IntelliVue MX450 Patient Monitor | Adverse Event Without Identified Device or Use Problem (2993) | Heart Failure (2206) | 2019-11-14 | Injury | A follow up report will be submitted once the investigation is complete | The customer reported a heart failure. Unknown, no further information is available at the time of the initial report
Philips Medical Systems | IntelliVue Information Center IX | Adverse Event Without Identified Device or Use Problem (2993) | ST Segment Elevation (2059) | 2019-09-19 | Injury | A follow up report will be submitted once the investigation is complete. Patient information requested not available at time of report. (b)(6). | The customer reported that when they do a print out from the bedside monitor mp30 and then the print out from the central (pic ix) for ecg, the printing is off by a 1-1. 5 ml. The ecg from the central monitor done at 23:09 showed significant st elevation in leads v2 v3 v4. Two minutes later at 23:11, they performed bed side generated ecg, which showed on st elevation, but the patient was thrombolysed based on the initial ecg. Patient was required to be flown to another location. The patient was administered the wrong medication. No information was received indicating the device was not a contributing factor in the event.
Philips Medical Systems | IntelliVue Information Center IX | Output Problem (3005) | No Code Available (3191) | 2018-10-23 | Injury | A follow up report will be submitted once the investigation is complete. Patient information requested, not available at time of report. | The customer reported a patient incident and indicated that if their equipment ¿compass¿, which provides the name associated with the antenna a unit is connected to, was working, they would have been able to attend the patient more quickly. The customer indicated a patient cardiac event.
Philips Medical Systems | IntelliVue MX500 Patient Monitor | Insufficient Information (3190) | Necrosis (1971) | 2017-09-16 | Injury | A follow up report will be submitted once the investigation is complete. | The customer reported necrosis/ signs of necrosis of 4 toes of the patient after non-invasive blood pressure measurement. The device was used for monitoring at the time of the alleged malfunction. Necrosis/ signs of necrosis of 4 toes. The patient is a few day old premature baby.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Incorrect Measurement (1383); Incorrect Or Inadequate Test Results (2456) | Cardiac Arrest (1762) | 2017-10-18 | Injury | A follow up report will be submitted upon completion of the investigation. Patient information was requested, unavailable at the time of this report. | The customer reported that "on (b)(6) 2017 between 14:59 and 15:28, 7 ecg measurements were taken on a patient, 5 ecgs were obtained with the intellivue multi measurement server x2 and 2 measurements were taken with an ecg cart. Different clinical symptoms were indicated throughout this time frame, and the customer observed inconsistencies in the 12-lead ecg taken with the philips device when compared to the ecg cart". The monitored patient suffered a cardiac arrest and was transferred to a different hospital. No information was provided on the measures taken to assist the patient and the current status of the patient.
Philips Medical Systems | IntelliVue Information Center IX | No Audible Alarm (1019) | Discomfort (2330) | 2017-10-03 | Injury | A follow-up report will be submitted once the investigation is complete. Patient information has been requested, but not available at the time of this report. | The customer reported that the patient in room 220 went into a desat condition at 8:23 am on (b)(6) 2017. The customer claims there was no audible alarm at the information center ix. The customer has stated that an "unfavorable outcome" occurred for this patient
Philips Medical Systems | IntelliVue MX700 Patient Monitor | Device Alarm System (1012) | No Code Available (3191) | 2017-08-17 | Injury | A follow up report will be submitted upon completion of the investigation. | The customer reported that "an incident where a monitored patient fell into a critical condition and no central alarm occurred". The device was used for monitoring at the time of the alleged malfunction. An adult male patient suffered an asystole. The patient was successfully reanimated.
Philips Medical Systems | IntelliVue MX800 Patient Monitor | No Audible Alarm (1019) | Cardiopulmonary Arrest (1765); Cyanosis (1798); Respiratory Distress (2045) | 2017-06-20 | Injury | Patient information was requested, unavailable at the time of this report. A follow up report will be submitted after philips obtains more information concerning this even | The customer reported that "the alarm was not audible" for room micu 960 on (b)(6)2017 at approximately 9:25 a. M. The device was used for monitoring at the time of the alleged malfunction. The patient was reported to be in distress and cyanotic and needed to be resuscitated.
Philips Medical Systems | IntelliVue Information Center IX | Device Alarm System (1012) | Low Oxygen Saturation (2477) | 2017-06-05 | Injury | A follow up report will be submitted upon completion of the investigation. Patient information was requested, unavailable at the time of this report. | The customer reported that the patient in room (b)(6) had a desat alarm on (b)(6) 2017 around 15:11 - 15:15, but the pop-up alarm failed to occur. The customer did note that the central station and bedside alarmed which a nurse happened to hear. It is indicated that the patient may have required emergent care, thus this will be considered a serious injury at this time.
Philips Medical Systems | IntelliVue Information Center IX | Computer Software Problem (1112); Defective Component (2292) | No Clinical Signs, Symptoms or Conditions (4582) | 2021-01-23 | Malfunction | A follow up report will be submitted once the investigation is complete. Phone number: (b)(6). | The customer reported that their loudspeaker was defective and required a software update. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Info Center IX | Unable to Obtain Readings (1516) | No Clinical Signs, Symptoms or Conditions (4582) | 2021-01-22 | Malfunction | A follow up report will be submitted once the investigation is complete | The customer reported that their patient information center (pic) was not reporting any ecg pauses. No patient involvement.
Philips Medical Systems | IntelliVue Information Center IX | No Audible Alarm (1019) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-28 | Malfunction | A follow up report will be submitted once the investigation is complete | The customer reported that their central station did not alarm for an spo2 issue that they believe should have caused an alarm. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | No Audible Prompt/Feedback (2282) | Insufficient Information (4580) | 2021-01-22 | Malfunction | Submission of a report does not constitute an admission that medical personnel, user facility, importer, distributor, manufacturer, or product caused or contributed to the event. | The customer reported that they had a speaker malfunction and no sound was audible. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue MP5 | Display or Visual Feedback Problem (1184) | Insufficient Information (4580) | 2020-12-24 | Malfunction | This is a supplemental report to mfr. Report #:1218950-2020-06242. The fda registration number initially chosen was incorrect. Please use this report. | The customer reported that the parameters are oscillating displaying saturation data not compatible with the clinical/hemodynamic stability of the patient. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Info Center IX | No Audible Alarm (1019) | No Clinical Signs, Symptoms or Conditions (4582) | 2021-01-20 | Malfunction | A follow-up report will be submitted once the investigation is complete. | The customer reported that they were not getting any audible alarms at their patient information center (pic). The device was in use on a patient. There was no report of patient or user harm.
Philips Medical Systems | IntelliVue X3 | No Audible Prompt/Feedback (2282) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-17 | Malfunction | A follow up report will be submitted once the investigation is complete | The customer reported that they had a speaker malfunction, there was no sound was audible. The device was not used for patient monitoring at the time of the alleged malfunction.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Output Problem (3005) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-17 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported spo2 is lower than expected measurements. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Information Center | No Audible Alarm (1019) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-02 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported that their central station was not audibly alarming. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Information Center IX | No Audible Alarm (1019) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-02 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported that their central station was not audibly alarming. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Audible Prompt/Feedback Problem (4020) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-11-29 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported a speaker volume issue on the intellivue multi measurement server x2. No information was provided whether any sound was still coming from the device. The device was not reported to be in use on a patient and no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Failure to Power Up (1476); No Audible Prompt/Feedback (2282) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-10-14 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported an inability to turn on even with a new battery. It was further stated that sound was not coming from the device. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue NMT Module | Incorrect, Inadequate or Imprecise Resultor Readings (1535); Failure to Calibrate (2440) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-10-08 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer called philips to report difficulties with nmt module calibration, and if the modules do calibrate, inaccurate readings have been observed. The device was reported to be in use on a patient, but no adverse event to patient or user was reported.
Philips Medical Systems | IntelliVue MX800 Patient Monitor | Display or Visual Feedback Problem (1184); No Audible Prompt/Feedback (2282) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-10-27 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported that a speaker malfunction inop was displayed on the intellivue mx800 patient monitor and no sound was coming from the device. No patient involvement.
Philips Medical Systems | IntelliVue MX500 Patient Monitor | Defective Component (2292); Device Displays Incorrect Message (2591) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-11-12 | Malfunction | A follow up report will be submitted once the investigation is complete. | The customer reported that a speaker malfunction inop was displayed on the intellivue mx500 patient monitor. No information was provided whether sound was still coming from the device. No patient involvement.
Philips Medical Systems | IntelliVue Multi Measurement Server X2 | Inaudible or Unclear Audible Prompt/Feedback (2283) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-11-11 | Malfunction | Initial reporter phone #: (b)(6). | The customer reported a speaker malfunction and there was no sound coming from the device. No patient involvement.
Philips Medical Systems | IntelliVue Information Center IX | Therapeutic or Diagnostic Output Failure (3023) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-11-09 | Malfunction | The reported malfunction was confirmed. Bed-to-bed alarm pop-up windows were not appearing. The remote service engineer (rse) was unable to determine what caused the issue to occur. The rse resolved the issue by rebooting the customers system. No further investigation is warranted. The device remains on site and in use. | The customer reported that the bed-to-bed alarm pop-up windows were not appearing. The device was in use on a patient. There was no report of patient or user harm.
Philips Medical Systems | IntelliVue MX40 802.11A/B/G | No Audible Prompt/Feedback (2282) | No Clinical Signs, Symptoms or Conditions (4582) | 2020-12-01 | Malfunction | A follow-up report will be submitted once the investigation is complete | The customer reported that the mx40 speaker does not work anymore. The device was not in use on a patient at the time of event