| Project: |  |  |  |
| --- | --- | --- | --- |
| CE Plan Title: |  |  |  |
| Owner: |  | Signature: |  |
| DHF Number: |  | Date: |  |

| Revision | Change Description in Brief |
| --- | --- |

| Approver |  |  |  |
| --- | --- | --- | --- |
| Function/Role | Name | Approver Signature | Date |

# Revision History

| Revision History (previous revisions) |  |
| --- | --- |
| Revision | Change Description in Brief |
| A.1 | Initial version |
| A.2 | Update before RfCuPT |
| A.3 | Added logistics |
| current revision see cover page |  |

Table of Contents

Revision History	2

1.	Summary	5

1.1	Objectives	5

2.	Device Description	5

2.1	Intended Purpose	5

2.2	Identification and Description of the Device	7

2.3	Clinical Development Plan	9

2.4	Clinical Background	10

2.5	New Features	12

2.6	Claims for the Device	12

3.	Scope	12

3.1	Manufacturer	12

3.2	Classification according to MDD / MDR	13

3.3	Risk Management:	14

3.4	Safety and Performance Claims and Potential Side-effects	14

4.	Equivalent Device	15

4.1	Conclusion on Equivalency	19

5.	Clinical Investigation (if applicable)	20

6.	Clinical Evaluation Strategy	20

6.1	Overview	20

6.2	Feature	21

7.	Updating the Clinical Evaluation Plan	23

8.	Analysis of the clinical data	23

9.	Writing the Clinical Evaluation report	23

10.	Choice of evaluators (qualification and conflict of interest)	23

11.	Referenced / Associated Documents	24

12.	Document Information	24

13.	Appendix A - Choice of literature and Conduct of the Study	25

13.1	Data Sources	25

13.2	Database search strategy	25

13.3	Exclusion criteria	28

13.4	Data Selection process	28

13.5	Clinical Data Appraisal Methods (based on MedDev 2.7.1. Rev 3)	28

13.6	Adverse Events	29

13.7	Data Appraisal Plan	30

13.8	Data reporting	35

14.	Appendix B – Plan for the Fulfillment of GSPR (General Safety and Performance Requirements)	36

# Summary

The Avalon fetal monitor & accessories will be evaluated through Literature Review, review of XXXX  held clinical data, Clinical Investigation(s) in the timeframe from initial release of Avalon FM20/30 in 2005 to the latest product enhancements up to SW release L.3 including the Avalon CL Fetal & Maternal Pod and electrode patch in 2018 & PMCF Studies.

## Objectives

The aim of clinical evaluation is to confirm clinical safety and performance by demonstrating conformity of Avalon Fetal/Maternal monitor family & accessories with the applicable Essential Requirements (MDD) or General Safety and Performance Requirements (EU MDR).

The objective of the Clinical Evaluation Plan is to describe the collection, assessment and analysis of the data regarding the clinical performance and safety of Avalon Fetal/Maternal monitor family & accessories and to document this against the clinical state of the art for the intended use of Avalon Fetal/Maternal monitor family & accessories in patients who are pregnant & fetuses.

# Device Description

## Intended Purpose

The XXXX  Avalon fetal/Maternal monitor family and accessories are diagnostic devices for pregnant women and fetuses.

### Intended Use and Patient Population

The XXXX  Avalon FM20 (M2702A), FM30 (M2703A), FM40 (M2704A) and FM50 (M2705A) Fetal/Maternal Monitors are intended for:

- Noninvasive monitoring of fetal heart rates and movements

- Noninvasive monitoring of maternal heart rates, maternal pulse rates, uterine activity, maternal noninvasive blood pressure, maternal oxygen saturation, and maternal temperature.

- Invasive monitoring of direct fetal heart rate, intrauterine pressure, and for displaying and recording of fetal and maternal electrocardiogram (ECG) (FM30 and FM50 only).

- Displaying, storing, and recording patient data and parameter values, and for generating alarms from fetal and maternal parameters.

- Transmitting patient data and parameter values to a patient information and surveillance system.

- Use by trained health care professionals.

- Use in antepartum testing areas, in labor and delivery rooms, and during postpartum recovery in the hospital environment. They are not intended for use in intensive care units or operating rooms.

- Monitoring in a bath or shower (Avalon CL cableless transducers Toco+, Toco MP, ultrasound, and ECG/IUP only).

- Transport situations in healthcare facilities, for healthcare facilities outside hospitals, such as doctors’ offices, and for use in private households (FM20 and FM30 only).

Patient Population:

The Avalon Fetal /Maternal Monitors and accessories is used for women who are at Gestational age >32 Weeks, with single fetuses, twins’ fetuses, and triplets’ fetuses.

The Avalon CL Fetal & Maternal Pod is used for women who are at >36 completed weeks, with singleton pregnancies.

### Clinical Environment

The Avalon Fetal/Maternal Monitors family and accessories are indicated for use by professionals whenever there is a need for monitoring the physiological parameters uterine activity, heart rate (Fetal/Maternal), electrocardiography (ECG), oxygen saturation, noninvasive blood pressure, and pulse rate, and temperature of pregnant women, and the fetal heart rates of single fetuses, twins, and triplets in labor and delivery rooms, in antepartum testing areas, in private households and during transports in healthcare facilities.

### Intended Users

According to the intended use the Avalon family of Fetal/Maternal monitors is to be used by trained health care professionals. Depending on local regulations and/or hospital policies these are registered nurses, gynecologists, midwives and health care assistants. In most geographies’ user training has to be documented.

### Intended Purpose, Warnings, Contraindications and Precautions

      - Intended Purpose

The XXXX  Avalon fetal/Maternal monitor family and accessories are diagnostic devices for pregnant women and fetuses.

      - Warnings & Contraindications

The Avalon Fetal/Maternal monitors and their accessories are not intended for use in an MRI environment or in an oxygen-enriched environment (for example, hyperbaric chambers). For Warnings related to Avalon Fetal/Maternal Monitors family and Accessories Please Refer the IFU (i.e. Avalon Fetal Monitor FM20/30, FM40/50, Avalon CL Release L3 With Software Revision L.3x.xx)

      - Precautions

The Avalon Fetal/Maternal Monitors are only for use on one patient at a time. They are not intended for home use.

Not therapeutic devices. The Fetal/Maternal monitors are for prescription use only.

Other General Precautions are mentioned below (IFU “Avalon Fetal Monitor FM 20/30, FM 40/50, Avalon CL Release L.3 with Software Revision L.3x.xx”):

- Always follow carefully and retain the instructions that accompany the specific cleaning and disinfecting substances you are using.

- Always dilute cleaning agents according to the manufacturer's instructions or use lowest possible concentration.

- Do not allow liquid to enter the case.

- Do not immerse the monitor in liquid. Protect it against water sprays or splashes.

- Do not pour liquid onto the system.

- Never use abrasive material (such as steel wool or silver polish).

- Never use bleach.

## Identification and Description of the Device

### Names and Product numbers

Avalon family product with accessories and part number are described in following Table 1.

Table 1: Avalon monitor and accessories with part number.

|  | Descriptions | Part Numbers |
| --- | --- | --- |
| Avalon Fetal / Maternal Monitors | FM20 | M2702A |
|  | FM30 | M2703A |
|  | FM40 | M2704A |
|  | FM50 | M2705A |
| Wired transducers / modules | Toco MP transducer | 867248, M2734B |
|  | Toco+ transducer | 867249, M2735A |
|  | Toco+ MP transducer | 867245 |
|  | Ultrasound transducer | 867246, M2736A, (M2736AA) |
|  | ECG/IUP transducer | 867247 |
|  | Remote Event Marker | 989803143411 |
| Cableless fetal transducers | Avalon CL base station | 866074 |
|  | Avalon CL Toco+ MP transducer | 866075 |
|  | Avalon CL US transducer | 866076 |
|  | Avalon CL ECG/IUP transducer | 866077 |
|  | Avalon CL Fetal & Maternal Pod | 866488 |
|  | Avalon CL Fetal & Maternal Patch (kit of 10) | 989803196341 |
| External devices | External power supply for FM20/30 *) | M8023A #E25 |
|  | Avalon CL Wide Range Pod | 866487 |
| Accessories | Reusable IUP Cable (Avalon) | 989803143931 |
|  | MECG ADAPT.CABLE | M1363A |
|  | Fetal monitoring recording paper | M1913A/ M1911A/ M1910A |
|  | IntelliVue CL Transmitter | 865221 |
|  | IntelliVue CL Charging Station | 865220 |
|  | X2/MP2 External Power Supply | M8023A |

### Description of the device and its intended application

The legally marketed devices XXXX  Avalon Fetal/Maternal Monitors FM20, FM30, FM40, and FM50 offer monitoring of fetal and maternal heart rates, uterine activity, maternal ECG wave, maternal noninvasive blood pressure (NIBP), oxygen saturation (SpO2) with pulse rate, and tympanic temperature during antepartum testing and labor and delivery.

FM20 and FM30 have a compact design and are operated from the top via a tiltable touch screen (Figure 1). FM40 and FM50 have a more traditional design and are operated from the front panel (Figure 2).

| Figure 1: Avalon FM20/30 | Figure 2: Avalon FM40/50 |
| --- | --- |
| Size: (W x H x D): 29cm x14cm x 34cm Weight: 5 kg | Size: (W x H x D): 43 cm x 17 cm x 37 cm Weight: 9 kg |

All Avalon Fetal/Maternal Monitors provide the following external measurement parameters:

- Up to three Fetal Heart Rates (FHR) via ultrasound (US)

- Fetal Movement Profile

- Uterine activity via external Toco

- Maternal Pulse Rate

- Maternal Heart Rate via maternal ECG

- Non-invasive blood pressure (NIBP)

- Maternal Pulse oximetry (SpO2)

- Maternal tympanic temperature

The Avalon FM30 is equal to the FM20 but offers additional parameters. The same applies to the FM50 when compared with the FM40.

Additional measurement parameters of FM30 and FM50;

- One Fetal Heart Rate and one fetal DECG wave via direct ECG (DECG), using a fetal scalp electrode

- Uterine activity via intrauterine pressure (IUP)

- Maternal ECG (MECG) wave

FM20 and FM30 optionally can be operated with a battery thereby allowing in-house transport.

All four Avalon monitors can be used together with the Avalon CL transducer system.

Figure 3: Avalon CL base station with docked transducers and an IntelliVue CL NBP pod.

Figure 4: CL Fetal & Maternal Pod and Beltless Solution

Avalon Fetal Monitors accessories are:

- Remote Event Marker

- Reusable IUP Cable (Avalon)

- MECG ADAPT.CABLE

- Fetal monitoring recording paper

- Avalon Toco+ MP Transducer

- Avalon ECG/IUP Transducer

- Avalon TOCO MP Transducer

- Avalon TOCO+ Transducer

## Clinical Development Plan

The clinical development for Avalon Fetal Monitors is described in following table.

Table 2: Avalon fetal monitor clinical development.

| Release | Release Content |
| --- | --- |
| 2005 | The Avalon FM20 and FM30 have been developed and brought to market. They succeed the Hewlett-Packard/Agilent/XXXX Series 50 fetal monitor family which first member was introduced in 1990. New features: - Graphical user interface with touch screen operation, leveraged from the XXXX IntelliVue patient monitors - Triplets monitoring via ultrasound - Toco transducer with almost flat surface to increase patient comfort The clinical evaluation was focusing on new or changed features, i.e. graphical user interface and the modified Toco transducer bottom. The triplet’s capability is an extension of the well-known twins monitoring. Triplets monitoring was possible with predecessor devices but required the combination of two fetal monitors. |
| 2008 | The Avalon FM40 and FM50 Fetal/Maternal monitors have been released. Intended use and core technology are the same as with FM20/30, but the industrial design fits additional clinical settings, especially when the monitor is mounted in a cart or cabinet requiring front panel operation. Touch display and display contents are identical to FM20/30. The front panel operation incl. recorder is well established since decades (e.g. Series 50 XM = M1350B). A detailed clinical evaluation therefore was not seen required for the changed housing. |
| 2010 | The following features were added: - FM20/30 battery operation for in-house transport - Toco MP transducer with integrated maternal pulse measurement from the abdominal skin - NST report following NICHD guidelines |
| 2014 | The Avalon CL cableless transducer system has been released, which is plugged into an Avalon fetal monitor. The functionality is widely based on the predecessor system “Avalon CTS” (M2720A) which had been released in 2003. New features of Avalon CL: - Cableless triplets monitoring via ultrasound - Cableless Toco MP and IUP measurements - Compatible to IntelliVue CL NBP and SpO2 pods for maternal surveillance, i.e. these pods can be assigned and charged on the Avalon CL base station. |
| 2018 | Software Rev L.3 with main features - Integration of the Monica Healthcare technology (“Helene”): fetal/ maternal heart rate and uterine activity through abdominal electrodes - Wide Range Pod (“Rhea”): connects the Saturn devices SpO2, NBP, F&M, and the Calimero CL Transducers US, Toco+ MP and ECG/IUP to a Calimero family fetal monitor - Integration of the trace interpretation according Dawes &Redman, addressing the market needs in EMEA and A/P |
| 2019 | The new Avalon Fetal Monitor Transducers have been released. The introduction of the 867xxx transducers replace the series with the 866xxx numbers. |

## Clinical Background

When a pregnancy occurs the healthcare provider will supply obstetrical care throughout the pregnancy which may include multiple visits to a physician office, clinic or hospital for evaluation. During a patient visit the mother is assessed and the well-being of the fetus is critically evaluated. The documentation of the maternal/fetal assessment (antepartum) is central to the accurate ongoing medical treatment of mother and child.

Auscultation of the fetal heart rate (FHR) became part of routine intrapartum care in many countries during the 19th century, and remains an important form of fetal surveillance, particularly in low-risk pregnancies and in low-resource countries. Several technical breakthroughs that occurred in the 20th century led to the development of different forms of continuous electronic monitoring of the FHR and uterine contractions in the 1950s and early 1960s. In late 1960s Hewlett-Packard (now XXXX ) introduced the first cardiotocography (CTG) device into the market. The term ‘Cardiotocography’ (kardia meaning heart, tokos meaning labor/childbirth) best describes the continuous monitoring of the FHR and uterine contractions, but other designations such as electronic fetal monitoring (EFM) are used as well.

Physiological parameters are usually measured by a Doppler ultrasound device strapped to the mother’s abdomen which detects fetal heartbeats via external monitoring. A tocodynamometer strapped on a second belt can detect changes in abdominal contour to identify contractions. The same data may be acquired internally by using a scalp electrode (DECG). A pressure sensor measures uterine activity.

In obstetrics, the interpretation of CTG or the failure to react in time is the major problem (intrapartum). The sooner the clinician knows about changes in fetal condition the more time he has for optimum patient care. Basic alerting informs when fetal tachycardia or bradycardia limit is exceeded during specified period of time, loss of signals, coincident detection of two fetal rates or fetal and maternal great rates. Advanced alerting features can analyze CTG patterns like baseline variability, accelerations, decelerations etc. and is based on fetal trace evaluations criteria defined by many leading obstetricians.

There are various Guidelines dealing with the importance of the electronic fetal monitoring and its interpretation and documentation:

Following medical societies have been chosen for a review of their guidelines:

  - RCOG (UK): The RCOG has been founded in 1929, has 16.000 members and is awarded by the UK National Institute for Health and Care Excellence (NICE) with the task of guideline development.

  - The Journal of Obstetrics and Gynaecology Canada (JOGC) publishes original research, reviews, case reports, and commentaries by Canadian and international authors, pertinent to readers in Canada and around the world. The Journal covers a wide range of topics in obstetrics and gynaecology and women’s health covering all life stages including the evidence-based Clinical Practice Guidelines, Committee Opinions, and Policy Statements that derive from standing or ad hoc committees of The Society of Obstetricians and Gynaecologists of Canada. The SOGC issued following documents:

  - NICHD is the U.S. government backed research organization for child health, including research in the perinatal field. The U.S. represent the biggest market for the XXXX  fetal monitors (ca 50%) and the biggest health market worldwide.

  - Founded in 1951 the American College of Obstetricians and Gynecologists (ACOG) is a membership organization dedicated to the advancement of women’s health care and the professional and socioeconomic interests of its members through continuing medical education, practice, research, and advocacy. It counts more than 58.000 members.

- Established in 1954 FIGO (International Federation in Gynecology and Obstetrics) counts member societies in about 130 countries or territories. It is the only global organization representing national societies of obstetricians and gynecologists. FIGO is in official relations with the World Health Organization (WHO) and a consultative status with the United Nations (UN).

  - The Association of Women's Health, Obstetric and Neonatal Nurses (AWHONN) established in 1969 is a nonprofit membership organization which represents 350,000 nurses in the U.S. who all share the same goal to promote the health of women and newborns.

The above recommendations and guidelines are fully supported by the various functions of the Avalon Fetal/Maternal monitors in all phases of care (ante, - intra and post-partum).

With Avalon Rev L the functionality has been further extended through the integration of the assessment of the fetal & maternal/ abdominal heart rate and uterine activity through abdominal electrodes, a technology provided by Monica Healthcare, UK, as well as the trace interpretation according Dawes &Redman.

The Avalon Fetal/Maternal monitors now in Rev L have been evolved through many releases (see chapter 2.3 above), backed by clinical evaluations either in situ with users and/or patient involvement or by selecting the clinical literature or clinical experience route. As each new release was an increment in functionality (either in new functions or just enhancements of already existing functions) the clinical evaluations followed the same incremental path as well.

## New Features

Calimero CL+ enhances the cableless Avalon CL solution in different dimensions as mentioned below:

- With the wide range option Avalon CL devices can be used in a hospital wide area leveraging the hospital’s WiFi installation.

- The fetal wireless patch solution leverages the Avalon CL concept and enhances the platform with a new measurement method.

In addition, the trace interpretation according Dawes and Redman leverages the existing implementation of the trace interpretation according NICHD guidelines, but addresses the market needs in EMEA and A/P.

## Claims for the Device

The Avalon fetal monitors are diagnostic devices for pregnant women and fetuses. There are no specific safety or performance claims.

# Scope

Scope of this Clinical Evaluation Plan is to:

- Identify the Essential Requirements (MDD) or General Safety and Performance Requirements (EU MDR) that need support from clinical data, (Refer Appendix B)

- Design a Clinical Evaluation strategy,

The Avalon Fetal/Maternal monitor family & accessories Clinical Evaluation Plan covers additional functionality of the Avalon Fetal/Maternal Monitors under development. It also clinically re- evaluates the Avalon Fetal /Maternal monitors with new Information’s from Literature and product Vigilance (i.e. PMCF studies) in the timeframe from 2005 to 2019 to enable a re-assessment of the benefit risk ratio.

## Manufacturer

XXXX  Medizin Systeme Böblingen GmbH
Hewlett-Packard-Strasse 2
71034 Böblingen, Germany

## Classification according to MDD / MDR

The XXXX  Avalon Fetal/Maternal Monitor Family FM20/FM30, FM40/ FM50, transducers, IUP catheter and Maternal Pods are classified as Class IIb devices and accessories as Class I devices as mentioned in the Table 3.

Table 3: Classification of Avalon Fetal/Maternal Monitor Family Devices According to MDD/MDR

| eDMR Part No. | Description | EU MDR Class |
| --- | --- | --- |
| 989803143411 | Remote Event Marker | I |
| 989803143931 | Reusable IUP Cable (Avalon) | I |
| M1363A | MECG ADAPT.CABLE | I |
| M1333A | Intrauterine Sensor-Tip Press. Catheter (OBL) | IIb |
| M2705A | Avalon FM50 Fetal Monitor | IIb |
| M2704A | Avalon FM40 Fetal Monitor | IIb |
| M2703A | Avalon FM30 Fetal Monitor | IIb |
| M2702A | Avalon FM20 Fetal Monitor | IIb |
| 866488 | Avalon CL Fetal & Maternal Pod | IIb |
| 866074 | Avalon CL Base station | IIb |
| M1913A/ M1911A/ M1910A | Fetal monitoring recording paper | I |
| 867245 | Avalon Toco+ MP Transducer | IIb |
| 867247 | Avalon ECG/IUP Transducer | IIb |
| 867248 | Avalon TOCO MP Transducer | IIb |
| 867249 | Avalon TOCO+ Transducer | IIb |
| 866075 | Avalon CL Toco + MP Transducer | IIb |
| 866076 | Avalon CL Ultrasound Tansducer | IIb |
| 866077 | Avalon CL ECG/IUP Transducer | IIb |
| 867246 | Avalon Ultrasound Transducer | IIb |
| 865220 | IntelliVue CL Charging Station | I |
| 865221 | IntelliVue CL Transmitter | I |

## Risk Management

The Safety Risk Assessment process complies with ISO 14971 – Medical Devices-Application of risk management to medical devices.  The results of the safety risk assessment (Doc-No: B3-P35430-SRA_approval_A.04.05) will be used as input for the clinical evaluation:

- To identify design features that require a clinical evaluation

- To verify the effectiveness of Risk control measures through clinical data, if applicable

- The information for safety (e.g. warnings, cautions, contraindications) communicated to the user in order to avoid harm.

Any new risks found during the Clinical Evaluation and not already addressed in the safety risk assessment must be fed back to the safety risk assessment process.

## Safety and Performance Claims and Potential Side-effects

The Safety claims for the family of the Avalon fetal/Maternal Monitors family are defined in the respective Calimero Safety Risk Analysis, Doc-No: B3-P35430-SRA_approval_A.04.05 filed in the respective Design History File. This document contains

•	A full list of the intended uses for the monitors and every single parameter/functionality

•	The product characteristics affecting safety,

•	All risks, hazard scenarios, risk estimations (probability, severity) and risk mitigations

•	The essential performance characteristics,

•	The expected service life and,

•	Key components (if any).

The Avalon monitors contain software. The software risk class has been classified as ‘C’ according to IEC 62304.

With respect to safety, no risks are identified in the risk management process that require escalation to the clinical evaluation in order to support MDD essential requirements / GSPR ER#1 (safety) and ER#6 (risk benefit). However, during the clinical evaluation, special attention shall be payed to the side-effects mentioned in the IFU of Avalon Fetal Monitor. If new risks will be identified during the planned Clinical Evaluation, which are not stated in the current SRA, the results will be reported to the risk management team.

- The Instructions of Use “Avalon Fetal Monitor FM20/30, FM40/50, Avalon CL Release L3 with Software Revision L3x.xx”, Doc-No 453564659391 contains all functionality descriptions, safety precautions and performance claims for the family of the Avalon Fetal Monitors.

### Identification of CRS requirements needing Clinical Evaluation

With this SW Release, following requirements have been identified to be clinically evaluated:

| Requirement ID | Requirement |
| --- | --- |
| M.3.1 | The assignment of the Pod is done on the Avalon CL base station the same way as the IntelliVue CL Pods. |
| M.3.2 | The measurement specific procedures and user interface guidance are equivalent to the predicate device (Monica Healthcare Novii Patch Solution). |

# Equivalent Device

Equivalent device is already known for Avalon Fetal Monitor and relevant information are given in Table 4.

Table 4: Equivalent device for Avalon fetal monitor.

| The Medical device (family) name | Model | Manufacturer’s information |
| --- | --- | --- |
| Novii (MONICA PATCH fetal and maternal monitor) Wireless Pod & Patch System | K140862 | Monica Healthcare (UK) |
| Avalon CL Fetal & Maternal Pod & Patch |  | XXXX Healthcare |

The Clinical, Technical and Biological characteristics are mentioned in below Table 5.

Table 5: Clinical, Technical and Biological characteristics of Avalon fetal monitor in comparison with competitor device.

|  | Monica Healthcare Novii Wireless Pod & Patch System (K140862) | XXXX Healthcare Avalon CL Fetal & Maternal Pod & Patch |
| --- | --- | --- |
| Pod and Patch |  |  |
| Patch |  |  |
| Indications for use | The Monica Novii Pod is an intrapartum maternal-fetal monitor that non-invasively measures and displays fetal heart rate (FHR), uterine activity (UA) and maternal heart rate (MHR). The Novii Pod acquires and displays the FHR tracing from abdominal surface electrodes that pick up the fetal ECG (fECG) signal. Using the same surface electrodes, the Pod also acquires and displays the UA tracing from the uterine electromyography (EMG) signal and the MHR tracing from the maternal ECG signal (mECG). The Pod is indicated for use on women who are at term (>36 completed weeks), in labor, with singleton pregnancies, using surface electrodes on the maternal abdomen. The Novii Patch is an accessory to the Novii Pod that connects directly to the Novii Pod and contains the surface electrodes that attach to the abdomen. The Novii Interface is an accessory to the Novii Pod which provides a means of interfacing the wireless output of the Novii Pod to the transducer inputs of a CTG Fetal monitor. The Novii Interface enables signals collected by the Novii Pod to be printed and displayed on a CTG Fetal Monitor and sent on to a central network, if connected. | Same Same Same The legally marketed XXXX Avalon CL Base Station provides the interface between the wireless output of the Avalon CL Fetal & Maternal Pod and the XXXX Avalon Fetal Monitor. |
| Intended Use | The Novii Pod maternal-fetal monitor and its accessories are intended for use by healthcare professionals | Same |
| Measured parameters | - Fetal Heart Rate (FHR) - Maternal Heart Rate (MHR) - Uterine Activity (UA) | Same |
| Target Patient Population | Pregnant women (Intrapartum > 36 completed weeks for labour & delivery) | Same |
| Principles of operation | - FHR picked up from the electrical signal produced by the fetal heart by the Novii Patch & Pod - MHR picked up from the electrical signal produced by the maternal heart by the Novii Patch & Pod - UA picked up from the electrical signal produced by the uterus during contractions by the Novii Patch & Pod - Novii Interface receives wireless data and converts signal to analogue for input into CTG monitor | - Same - Same - Same - Legally marketed XXXX Avalon CL Base Station receives wireless data and feeds them into the XXXX Avalon Fetal Monitor |
| Measurement methods for FHR & MHR | Abdominal electrophysiological signal and template matching, filtering and confidence tagging to identify the FECG and MECG complexes from other signals and interference/noise. R-R interval detection to determine FHR & MHR | Same |
| Measurement methods for UA | UA picked up from the electrical signal on the abdomen produced by the uterus during contractions | Same |
| Parameters | FHR: Range 60-240 beats per minute Resolution: 1/4 BPM Produced 4 time per second from a rolling 2s average MHR: Range 40-240 beats per minute Resolution: 1/4 BPM Produced 4 time per second from a rolling 2s average UA: Range 0-500 microvolts Resolution: 0-255 levels representing 100% of full scale Produced 4 time per second from a rolling 2s average | Same (labeled “aFHR”) Same (labeled “aHR”) Same (labeled “aToco”) |
| Material in contact with human tissues | Monica Novii Patch | XXXX Avalon CL Fetal & Maternal Patch |
| Cleaning & Sterilization | The Novii Pod & Interface are reusable and require cleaning in according with manufacturer’s instructions. No need for sterilisation The Monica Novii Patch is single use | Same |
| Accessories | Monica Novii Patch and skin preparation tape | XXXX Avalon CL Fetal & Maternal Patch and skin preparation paper |
| Data display & format | Novii Interface interfaces to a standard CTG monitor with standard 1 or 2 or 3cm\min display for FHR\MHR& UA display | Avalon CL Base Station interfaces only the Avalon Fetal Monitor for FHR/MHR & UA display |
| Power supply | Novii Pod uses 3.3V Lithium Polymer battery Novii Interface uses mains adaptor | Avalon CL F&M Pod uses 3.7V Li-Ion battery Avalon CL Base Station uses the power from the connected Avalon Fetal Monitor |
| Monitor outputs | Novii Pod uses Series 50 format data over Bluetooth interface Novii Interface accepts Series 50 format data from the Novii Pod and outputs analogue data formatted for input to a CTG fetal monitor transducer inputs | Avalon CL Fetal & Maternal Pod uses SRR (Short Range Radio) Avalon CL Base Station receives data from the Pod via radio and provides them digitally to the Avalon Fetal Monitor, via the connection cable. |
| Shelf Life | The Novii Pod & Interface: No The Novii Patch: Yes | Same |
| Size & Weight | Novii Pod: 43 x 36 x 12 mm, 30g Novii Patch: 185 x 160 mm, 13g | Avalon CL F&M Pod: 63 x 20 x 49 mm, 72g Avalon CL F&M Patch: 161 x 9 x 190 mm, 20g |
| Operating temperature | 0 to + 30°C | +10°C to +40°C (Pod & Patch) |
| Storage temperature | 0 to + 30°C | -20°C to +60°C (Pod) +10°C to +30°C (Patch) |
| Data Interfaces (Pod) | Input: Electrical signals from the Patch Output: Wireless data transmission to the Novii Interface Wireless connection: Bluetooth Class 1 Operational Range: 100 meters (300 feet) – line of sight | Input: Same Output: Wireless data transmission to the legally marketed XXXX Avalon CL Base Station Wireless connection: Short Range Radio IEEE 802.15.4 Operational Range: in room 30 meters (90 feet) |
| Power (Pod) | Battery: Rechargeable lithium polymer 3.7V 750mA-hours Charging: Contactless charging with the Novii Interface (105-PT-001) | Battery: Rechargeable Lithium Ion 3.7V 1150mA-hours Charging: Charging with contacts at the legally marketed XXXX Avalon CL Base Station |
| Data conversion & encryption (Patch) | Analogue, converting electrophysiological signals picked up via electrode areas from ionic to electronic conduction Microchip containing a factory pre-set code to prevent counterfeiting measure and to facilitate pairing | Same |
| Ingress protection (Pod) | IPx0 | IP67 |
| Battery run time | >14 h | >16h |

## Conclusion on Equivalency

The equivalency will be checked for the device under evaluation with respect to the above equivalent device. Literature and other clinical data in the CER will support this conclusion.

The Novii patch and pod have been integrated into the Avalon Fetal/Maternal monitors as ‘Avalon CL Fetal & Maternal Pod’ and ‘Avalon CL Fetal & Maternal Patch’. The Avalon CL F&M Pod and Patch offer the same clinical features as the Novii patch and pod and have equivalent specifications.

# Clinical Investigation (if applicable)

The following accessories required no specific investigation because of technical equivalence (Table 6).

Table 6: Avalon fetal monitor clinical investigation details.

| Transducer name | Cableless / old | Cable / new |
| --- | --- | --- |
| Avalon Cableless Toco + MP Transducer | 866075 | 867245 |
| Avalon Cableless Ultrasound Tansducer | 866076 | 867246 |
| Avalon Cableless ECG/IUP Transducer | 866077 | 867247 |
| Avalon Toco MP Transducer |  | 867248 |
| Avalon TOCO + Transducer |  | 867249 |

# Clinical Evaluation Strategy

## Overview

Clinical evaluation for the Avalon Fetal/Maternal Monitor family and accessories will use literature review and a combination of different clinical data collection.

The intended literature review will be enhanced by a re-evaluation of earlier XXXX  owned clinical data (previous studies and PMS generated data) of previous releases of the Avalon fetal/maternal Monitor Family. Furthermore, the clinical evaluation report will include an updated search for complaints and patient incidents.

The device is well established and low risk. All available relevant data (both positive and negative) will be taken into account. Also, for all literature articles and investigations identified as relevant for this clinical evaluation, the appraisal and weighting of the articles and or study reports will be documented.

The contents of information materials (e.g. IFU, available promotional materials, including additional accompanying documents if relevant) will be verified during the clinical evaluation activities.

### Identification of Pertinent Data

The safety and performance profiles for the Avalon fetal monitors are well established and documented. No new design features, additional indications, which might necessitate a clinical investigation, have been incorporated in the devices; nor have any specific clinical concerns emerged which might warrant such.

Therefore, this clinical evaluation for Avalon fetal monitors will draw upon safety and performance from both non-clinical and clinical data sets as set out below.

•	Non-clinical data (i.e., the verification / validation testing data) from the device under
        evaluation)

•	Previous Clinical investigation data from the device under evaluation

•	Clinical literature data from the device under evaluation

•	Safety data from the device under evaluation

•	PMS data, including FSCAs, recalls, and internal complaints trending and analysis

•	Risk management data

## Feature – Avalon Fetal /Maternal Monitor Family L3

This Clinical Evaluation Plan refers to the features of L.3 of the Avalon Fetal/Maternal Monitor Family. Hence it will focus on the beltless solution

- Actual patch application (skin preparation, durability of electrode adhesive)

- The related user interface (e.g. impedance quality indication)

- Patient comfort while wearing Patch and Pod

### Methods

Methods are:

- Pre-Clinical data (i.e., Verification / Validation test results) from the device under evaluation

  - Performance and safety will be evaluated, separate from the clinical evaluation, in simulated environments (not in a clinical setting, not by clinical users).

  - XXXX  will use literature in order to demonstrate feasibility and clinical benefits of the technology. Those data will have been produced with the Monica devices Novii and AN24 which use same or similar technology.

- CE/CE, CE/LIT, CE/CI Clinical Experience; i.e. market experience, Post distribution surveillance, review of XXXX  owned clinical data, user panels, Scientific Literature Review, Clinical Investigation

  - Method of choice is CE/CE. XXXX  will use clinical data regarding user preferences, user interface and patient comfort collected with released devices in a clinical setting.

### Research Question

The research questions to be answered by the safety and performance literature search and review are below.

  - Research Question #1 (R-1): Were the Avalon fetal monitors mentioned in the literature evaluated to be generally safe and effective?

  - Research Question #2 (R-2): What are the risks and benefits of the Avalon fetal monitors?

The research question to be answered by the AE, recall, and FSCA database search and review is below.

- Complaints Criteria #1 (CC-1): Were risks or safety issues (i.e., adverse effects/events, recalls, field safety corrective actions) reported for patients for which the Avalon fetal monitors were used as indicated?

### Hypothesis

There is no hypothesis to be validated. Goal of the clinical evaluation is to collect user feedback for further product optimization. Subject of the clinical evaluation are the new features in release L.3 of the Avalon Fetal/Maternal Monitor Family and accessories.

### Timeframe / Schedule (planning, executing, reporting)

For clinical investigations the typical sequence is: identification of investigational site(s), CIP generation, IRB/EC submission, research collaboration agreement with site signed, IRB/EC approval, installation and training, study period with clinical data collection, statistical evaluation of results, CIR approved).

For Literature Review the typical sequence is: Identification of Search Protocol (see Appendix A of this plan), Appraisal Criteria Definition, Identification of Sources, Literature Search, Evaluation of Output Results (quality, fitness to search goals), Review and Analysis (suitability and contribution), conclusions;

For user panels the typical sequence is planning (site, invitations), generation of use scenarios and questionnaires, dry run, run the user panel, statistical evaluation of results, report.

### Resources

XXXX  infrastructure will be used for the clinical evaluation incl. literature review via common public databases as well as via search of internal complaint and non-compliant databases.

The clinical evaluation requires:

- 1 functional prototype of the Avalon CL Fetal & Maternal Pod

- 1 fetal Monitoring Electrode Patch

- 1 FM 30 host monitor with software compatible to the Avalon CL Fetal & Maternal Pod

- Monica CE Report (external, TFS# 138974)

The resources needed for the post-release evaluation are specified for each site in chapter 6.2.6 “Investigational sites”

### Investigational sites

The clinical evaluation shall involve users from several European hospitals from:

•	University Hospital Tübingen, Germany

•	University Hospital Marburg, Germany

### Logistics

Table 7: Logistic details of Avalon fetal monitor clinical investigation.

| Sl. No | Potential sites | Recourses |
| --- | --- | --- |
| 1 | Tübingen, University Hospital | - 3 Avalon CL Fetal & Maternal Pods - Avalon CL Fetal & Maternal Patches, as required (single use, disposable. Total number will depend on demand) - ongoing and started in 2019 |
| 2 | Marburg, University Hospital | - 3 Avalon CL Fetal & Maternal Pods - Avalon CL Fetal & Maternal Patches, as required (single use, disposable). Total number will depend on demand) - ongoing and started in 2019 |

# 6.3 Various other activities planned

(Like user panels, simulation sessions etc.)

# Updating the Clinical Evaluation Plan

The clinical evaluation plan will be updated where applicable throughout the complete product life cycle to maintain conformity with the relevant Essential Requirements and / or General Safety Performance Requirements (EU MDR requirements).

The frequency of active clinical evaluation updates and a justification will be documented in the clinical evaluation report.

# Analysis of the clinical data

The clinical data will be analyzed in the Clinical Evaluation Report with the aim to demonstrate conformity of Avalon Fetal/Maternal Monitor Family with the Essential Requirements (including ER1, ER3, and ER6) of the MDD and General Safety and Performance Requirements of MDR(Including GSPR 1, GSPR 5 and GSPR 8) on performance and safety of the device.

With respect to residual risks and uncertainties or unanswered questions (including on rare complications, long term performance, safety under wide-spread use), will be determined in the CER whether these: -

- Are acceptable for CE-marking,

- Are required to be addressed during Post Market Surveillance and

- With relevant Post Market Clinical Follow-up.

# Writing the Clinical Evaluation report

The Clinical Evaluation Report (CER) will summarize and draw together the evaluation of all the relevant clinical data documented or referenced in other parts of the technical documentation and the results of the clinical data analysis described in paragraph 8. The CER and the relevant clinical data constitute the clinical evidence for conformity assessment.

In the final CER it will be discussed if sufficient clinical evidence for demonstration of the clinical performance and clinical safety of the device under evaluation has been obtained and, whether there are adequate data for all aspects of the intended purpose and for all products/ models/ sizes/ settings covered by the Clinical Evaluation.

The CER will describe the benefits and risks of the device (their nature, probability, extent, duration and frequency); explain the acceptability of the benefit/risk profile according to current knowledge/ the state of the art in the medical fields concerned, with reference to applicable standards and guidance documents, available medical alternatives, and the analysis and conclusions of the evaluators on fulfilment of all Essential Requirements pertaining to clinical properties of the device (MDR , MDD ER1, ER3, ER6).

| Date | Responsibility | Name | Timeline |
| --- | --- | --- | --- |

# Choice of evaluators (qualification and conflict of interest)

| Name the clinical evaluators | Responsibility | Qualifications | Date | Signature |
| --- | --- | --- | --- | --- |
| Dr. Benedikt Latz |  | Dr. Benedikt Latz is a medical doctor. He is leading the clinical evaluation of XXXX fetal monitors for more than 10 years. His qualification is on file available on demand (e.g. personal file). |  |  |
| Rebecca Schiller |  |  |  |  |

# Referenced / Associated Documents

Following is the complete list of all documents and other sources of information referenced in the clinical evaluation Plan.

| Document Name | Document No. | Revision |
| --- | --- | --- |
| Calimero CL+ Customer Requirements Specifications (CRS) | B3-P35440-CRS | A.01.10 |
| Calimero Safety Risk Assessment | B3-P35430-SRA | A.04.05 |
| Calimero CL+ Verification and Validation Plan | B3-P35440-VV-Plan | XX |

# Document Information

References

| Document Number | Document Title | Type |
| --- | --- | --- |
| EN/ISO 14971 | Application of Risk Management to Medical Devices | E |
| 93/42/EEC | MDD: Directive Concerning Medical Devices | E |
| (EU) 2017/745 | MDR: Regulation (EU) of the European Parliament and of the Council on medical devices, and amending Directive 2001/83/EC, Regulation (EC) No 178/2002 and Regulation (EC) No 1223/2009 | E |
| MEDDEV 2.7/1 Rev 4 | Clinical Evaluation: a guide for manufacturers and notified bodies under directives 93/42/EEC and 90/385 EEC (MEDDEV 2.7/1 Rev 4) | E |
| EN ISO 14155:2011 | Clinical investigation of medical devices for human subjects – Good clinical practice | E |

# Appendix A - Choice of literature and Conduct of the Study

## Data Sources

| Literature database | Justification |
| --- | --- |
| PubMed/ PubMed Central | PubMed comprises over 22 million citations for biomedical literature from MEDLINE, life science journals, and online books. PubMed citations and abstracts include the fields of biomedicine and health, covering portions of the life sciences, behavioural sciences, chemical sciences, and bioengineering Link: “(https://www.ncbi.nlm.nih.gov/pubmed/advanced)” |
| Cochrane | The Cochrane Library (ISSN 1465-1858) is a collection of databases that contain different types of high-quality, independent evidence to inform healthcare decision-making. Link: “(http://onlinelibrary.wiley.com/cochranelibrary/search)” |

## Database search strategy

There will be 2 database search strategies:

- For safety and performance

- For the State of the Art

The results may be combined (especially if the literature research results coincide) in one appraisal section.

- Activated Limits:

- English

- full text available or/and abstract only

- publication date <10 years

- Search Fields: All Text Fields

- Search Terms:

- (“FM20” OR “FM30” OR “FM40” OR “FM50”) and XXXX : 0 results

- XXXX  AND “fetal monitor”: 3 results

- Principles AND “fetal monitoring”: 7 results

- Fetal AND monitoring AND Cochrane: 75 results

- Continuous cardiotocography: 44 results

- noninvasive AND “blood pressure” AND pregnancy: 6 results

- Fetal AND ECG:  234 results

- maternal AND pulse AND oximetry AND labour: 8 results

Search strings:

- Search with (((((FM20) OR FM30) OR FM40) OR FM50) AND XXXX ) AND ("2015/12/01"[Date - Publication]: "3000"[Date - Publication]) gathered 0 results

- Search with (XXXX [All Fields] AND (("fetus"[MeSH Terms] OR "fetus"[All Fields] OR "fetal"[All Fields]) AND ("monitoring, physiologic"[MeSH Terms] OR ("monitoring"[All Fields] AND "physiologic"[All Fields]) OR "physiologic monitoring"[All Fields] OR "monitor"[All Fields]))) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) gathered 1 result, non-relevant (related to gas monitoring full-term infants immediately after caesarean section)

- The search with ((principles[Title/Abstract] AND fetal[Title/Abstract]) AND monitoring[Title/Abstract]) AND ("2015/12/01"[MDAT] : "3000"[MDAT]) uncovered 1 result dealing with health information technology (IT) which has been shown to promote patient safety in Labor and Delivery (L&D) units (not relevant).

- Search with ((fetal[Title/Abstract] AND monitoring[Title/Abstract]) AND Cochrane[Title/Abstract]) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) retrieved 22 records; excluding reports dealing with  glycose monitoring, fertility, pharmacokinetics, umbilical ultrasound and dystocia resulted in 5 papers to be considered.

- Search with (continuous [All Fields] AND ("cardiotocography"[MeSH Terms] OR "cardiotocography"[All Fields])) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) retrieved 25 records. Excluding papers already retrieved by previous searches, clinical practice improvements not related to CTG, pharmacological interventions or research work on new CTG interpretation algorithms resulted in 6 papers considered for this review.

- Search with ((non-invasive[All Fields] AND ("pressure"[MeSH Terms] OR "pressure"[All Fields])) AND ("pregnancy"[MeSH Terms] OR "pregnancy"[All Fields])) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) resulted in 26 records; after excluding papers on pharmacokinetics, DNA screening, bilirubin measurements , echo cardiography and alternative NBP measurement methods to assess pre-eclampsia none remained for consideration.

- The search with (fetal[Title/Abstract] AND ECG[Title/Abstract]) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) resulted in 86 records; after excluding papers dealing with anesthesia during caesarian sections, physiology or new algorithm research, already retrieved articles and pharmacological studies 19 remained for consideration , 3 found relevant to fetal / maternal monitoring.

- The search with (((("mothers"[MeSH Terms] OR "mothers"[All Fields] OR "maternal"[All Fields]) AND ("pulse"[MeSH Terms] OR "pulse"[All Fields] OR "heart rate"[MeSH Terms] OR ("heart"[All Fields] AND "rate"[All Fields]) OR "heart rate"[All Fields])) AND ("oximetry"[MeSH Terms] OR "oximetry"[All Fields])) AND ("labour"[All Fields] OR "work"[MeSH Terms] OR "work"[All Fields] OR "labor"[All Fields] OR "labor, obstetric"[MeSH Terms] OR ("labor"[All Fields] AND "obstetric"[All Fields]) OR "obstetric labor"[All Fields])) AND ("2015/12/01"[PDAT] : "3000"[PDAT]) revealed 6 records, 3 have been considered relevant.

The following literature research covered the periods:

13-Feb-2018 until 25-Jul-2019 and 26-Jul-2019 until 31-Jan-2020

Literature sources used to identify data: PubMed and Cochrane

Type of search: online

Activated limits:

- Starting dates as shown above

- English

- Full text and abstract

| Database search details: # | Terms | Source |
| --- | --- | --- |
| Search A | Search with "FM20"[All Fields] OR "FM30"[All Fields] OR "FM40"[All Fields] OR "FM50"[All Fields] AND XXXX [All Fields] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search B | XXXX [All Fields] AND "fetal monitor"[All Fields] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search C | Principles[All Fields] AND "fetal monitoring"[All Fields] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search D | "Fetal monitoring"[All Fields] AND Cochrane[All Fields] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search E | "Continuous cardiotocography"[All Fields] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search F | noninvasive[All Fields] AND "blood pressure"[All Fields] AND ("pregnancy"[MeSH Terms] OR "pregnancy"[All Fields]) AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search G | "Fetal monitoring"[All Fields] AND ("electrocardiography"[MeSH Terms] OR "electrocardiography"[All Fields]) AND (hasabstract[text] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) AND English[lang]) | PubMed |
| Search H | ("pregnancy"[MeSH Terms] OR "pregnancy"[All Fields]) AND ("mothers"[MeSH Terms] OR "mothers"[All Fields] OR "maternal"[All Fields]) AND ("pulse"[MeSH Terms] OR "pulse"[All Fields] OR "heart rate"[MeSH Terms] OR ("heart"[All Fields] AND "rate"[All Fields]) OR "heart rate"[All Fields]) AND ("oximetry"[MeSH Terms] OR "oximetry"[All Fields]) AND (hasabstract[text] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT])) | PubMed |
| Search I | ("mothers"[MeSH Terms] OR "mothers"[All Fields] OR "maternal"[All Fields]) AND ("labour"[All Fields] OR "work"[MeSH Terms] OR "work"[All Fields] OR "labor"[All Fields] OR "labor, obstetric"[MeSH Terms] OR ("labor"[All Fields] AND "obstetric"[All Fields]) OR "obstetric labor"[All Fields]) AND ("pulse"[MeSH Terms] OR "pulse"[All Fields] OR "heart rate"[MeSH Terms] OR ("heart"[All Fields] AND "rate"[All Fields]) OR "heart rate"[All Fields]) AND ("oximetry"[MeSH Terms] OR "oximetry"[All Fields]) AND (hasabstract[text] AND ("2018/02/13"[PDAT] : "3000/12/31"[PDAT]) | PubMed |
| Search J | Cochrane Reviews matching "FM20" in Title Abstract Keyword AND "FM30" in Title Abstract Keyword AND "FM40" in Title Abstract Keyword AND "FM50" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search K | Cochrane Reviews matching "FM20" in Title Abstract Keyword AND "FM30" in Title Abstract Keyword AND "FM40" in Title Abstract Keyword AND "FM50" in Title Abstract Keyword AND XXXX in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search L | XXXX in Title Abstract Keyword AND "FM20" in Title Abstract Keyword OR "FM30" in Title Abstract Keyword OR "FM40" in Title Abstract Keyword OR "FM50" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search M | Principles in Title Abstract Keyword AND "fetal monitor" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search N | "Continuous Cardiotocography" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search O | noninvasive in Title Abstract Keyword AND "blood pressure" in Title Abstract Keyword AND pregnancy in Title Abstract Keyword OR "fetal monitoring" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched | Cochrane |
| Search P | "Fetal Monitoring" in Title Abstract Keyword AND electrocardiography in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search Q | pregnancy in Title Abstract Keyword AND maternal in Title Abstract Keyword AND monitoring in Title Abstract Keyword OR oximetry in Title Abstract Keyword AND "heart rate" in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Search R | obstetrics in Title Abstract Keyword AND cardiotocography in Title Abstract Keyword AND interpretation in Title Abstract Keyword - with Cochrane Library publication date Between Feb 2018 and Jul 2019 (Word variations have been searched) | Cochrane |
| Sum |  |  |
| Relevant Papers |  |  |

## Exclusion criteria

The total set of found literature will be filtered using the following criteria:

| EC # | Exclusion Criterion |
| --- | --- |
|  | Publications, focusing on monitoring and guidance |
|  | Publications, focusing on simulation |
|  | Publications related to non-human experience |
|  | Publications related to surgical procedures |
|  | Publications unrelated to patient monitoring |
|  | Paper does not relate to safety, performance or potential risk of the fetal monitors. |

## Data Selection process

Documents retrieved will be studied and weighted according to their significance. Only peer reviewed papers (no conference abstracts), published in the English language will be included. All identified papers will be assessed for relevance with respect to clinical, technical and biological parameters and their similarities to the device under consideration.

## Clinical Data Appraisal Methods (based on Meddev 2.7.1. Rev 4)

Appraisal Criteria for Suitability (identical to Table D1 of Appendix D of the MEDDEV 2.7.1)

| Suitability Criteria | Description | Grading System |  |
| --- | --- | --- | --- |
| Appropriate device | Were the data generated from the device in question? | D1 D2 D3 | Actual device Equivalent device Other device |
| Appropriate device application | Was the device used for the same intended use (e.g., methods of deployment, application, etc.)? | A1 A2 A3 | Same use Minor deviation Major deviation |
| Appropriate patient group | Where the data generated from a patient group that is representative of the intended treatment population e.g., age, sex, etc.) and clinical condition (i.e., disease, including state and severity)? | P1 P2 P3 | Applicable Limited Different population |
| Acceptable report/data collation | Do the reports or collations of data contain sufficient information to be able to undertake a rational and objective assessment? | R1 R2 R3 | High quality Minor deficiencies Insufficient information |

Note: Optimum Combination: D1 + A1 + P1 + R1

Appraisal Criteria for Data Contribution

| Contribution Criteria | Description | Grading System |  |
| --- | --- | --- | --- |
| Data source type | Is the functionality in question directly mentioned by the study/paper and/or was the design of the study appropriate? | T1 T2 | Yes No |
| Outcome measures | Do the outcome measures reported reflect the intended performance / intended use of the device? | O1 O2 | Yes No |
| Follow up | Is the duration of follow up long enough to assess whether duration of treatment effects and identify complication? | F1 F2 | Yes No |
| Statistical Significance | Do the data include any quantitative / qualitative statements? / Has statistical analysis been provided and is it appropriate | S1 S2 | Yes No |
| Clinical Usage / Clinical Significance | Was the device with the functionality in question used for day-to-day service and was clinically significant? | C1 C2 | Yes No |

Note: Optimum combination: T1 + O1 + F1 + S1 + C1

Note: Reason for selecting these criteria: Table D2 of Appendix D of the MEDDEV 2.7.1 is merely aiming on treatment results of therapeutic devices (e.g. patient outcome, long term treatment effects, and complications). The criteria below are adapted to the nature of the monitoring devices (with uninterrupted diagnostic functionality only).

## Adverse Events

Other data, including PMS data from adverse event databases maintained by various Competent Authorities, including the FDA MAUDE database, EUDAMED (when available), MHRA, BfArM, TGA, etc., will be sought.

| Database | Link | Comments |
| --- | --- | --- |
| FDA – Safety Communications | www.fda.gov/MedicalDevices/Safety/AlertsandNotices/default.htm | This database contains safety communications used to alert health professionals and the public about newly observed potential risks to ensure the proper use of medical devices and the health and safety of patients. |
| FDA MAUDE - Manufacturer and User Facility Device Experience | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/TextSearch.cfm | The MAUDE database houses medical device reports submitted to the FDA by mandatory reporters (manufacturers, importers and device user facilities) and voluntary reporters such as health care professionals, patients and consumers. |
| CDRA – Medical Device Recalls | http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm | FDA database. This database contains consumer information about recalls of medical devices for which a reasonable chance exists that they could cause serious health problems or death. |
| Health Canada – Recalls and safety alerts | www.healthycanadians.gc.ca/recall-alert-rappel-avis/index-eng.php | The Recalls and Safety Alerts web application provides easy access to a comprehensive list of recalls, advisories, and safety alerts. This web application includes recalls from Health Canada, the Canadian Food Inspection Agency, and Transport Canada. |
| MHRA – Alerts and recalls | www.gov.uk/drug-device-alerts | Alerts and recalls database for drugs and medical devices from the Medicines & Healthcare products Regulatory Agency, UK. |
| BfArM – Field Corrective Actions | www.bfarm.de/SiteGlobals/Forms/Suche/EN/kundeninfo_Filtersuche_Formular_en.html | German Competent Authority Field Corrective Action database. Field Corrective Actions include removals of medical devices from the market or any other corrective action on devices in use. |
| TGA Australia - Database of Adverse Event Notifications - medical devices | https://apps.tga.gov.au/prod/DEVICES/daen-report.aspx | Therapeutic Goods Administration (TGA) is part of the Australian Government Department of Health; The DAEN - medical devices contains information from reports of adverse events that the TGA has received in relation to medical devices used in Australia. |

- EUDAMED (when available)

- EMBASE (when available)

Adverse Events, Recall, FSCA

| ID# | Search strategy | Database |
| --- | --- | --- |
| Adverse Events search |  |  |
| SA1 | "Product Code: ""HGM""; Date Report Received by FDA: 13-Jul-2017 to 19-Feb-2018" | FDA Maude Database |
| SA2 | Manufacturer: Monica Healthcare Event Type: Death Report Date From: 01/01/2010 Report Date To: 01/31/2018 | FDA Maude Database |
| SA3 | Monica Healthcare Event Type: Injury Report Date From: 01/01/2010 Report Date To: 01/31/2018 | FDA Maude Database |
| SA4 | Manufacturer: Monica Healthcare Event Type: Malfunction Report Date From: 01/01/2010 Report Date To: 01/31/2018 | FDA Maude Database |
| Recall, FSCA |  |  |
| SR1 | "Product Code: ""HGM""; Recall Date: 13-Jul-2017 to 19-Feb-2018", performed on 19-Feb-2018 | Medical Device Recalls / FDA |

## Data Appraisal Plan

All collected clinical data will be appraised according to the rules mentioned above in 13.5.  This applies also to existing clinical data, e.g. from previous studies or CE Reports.

XXXX 
The following list contains all clinical evaluation reports or XXXX  initiated studies with Avalon fetal monitor. The respective TFS (Team Foundation Server database) entry contains an overview of the specific clinical evaluation method(s), a summary of the results and a full copy of the report.

| Report Title | Date | Functionality covered | TFS # | Summary of results |
| --- | --- | --- | --- | --- |
| Avalon FM user preference CE Report | 25-Jul-2005 | M2703A, “Calimini” | 58870 | The Avalon FM M2703A fetal monitor is fully usable within its intended use. This is valid for usability, patient comfort and clinical performance. |
| Caliretro CE Report | 26-Feb-2008 | M 2704/M2705A, “Caliretro” | 58871 | The Calimini M2705A fetal monitor is fully usable within its intended use. This is valid for usability, patient comfort and clinical performance. |
| Calimero HIF CE Report | 07-Jan-2004 | M 2703A “Calimini” | 89226 |  |
| Calimini (M2703A) Human Interface 2 Clinical Evaluation Report | 04-May-2005 | M 2703A “Calimini” | 89247 |  |
| Fetal Heart Rate (FHR) from Ultrasound (US) | 04-May-2005 | M 2736A Ultrasound Transducer | 89615 | Comparative testing with a representative test suite of ultrasound raw signals taken from a large data base with recorded patient signals demonstrated that the performance of the modified ultrasound transducer with SW Rev. A.06.31 is equivalent to the ultrasound parameter of Series 50. It was also demonstrated that the performance of the Avalon transducer with SW Rev. A.06.31 is better than its predecessor SW Rev. A.04.30. |
| Calimero CL – Clinical Evaluation Report | 15-May-2014 | Avalon Cableless Ultrasound Transducer Product no: 866076 | 89613 |  |
| Fetal Movement Profile (FMP) | 01-Jul-2014 | M2702/-05 “Avalon fetal monitors” SW Rev J.3 onwards | 89614 |  |
| Fetal Heart Rate from DECG | 26-Oct-2003 | 989803137631, fetal scalp electrode for DECG measurement | 89613 | - Personal control of application procedure: All midwifes said it was adequate or good (57%). No injury has been noted on any patient. - Perceived ease of application: All midwifes but one said it was acceptable or easy to use (64%). - Signal quality: The fetal scalp electrode developed by Clinical Innovations Inc, Utah, USA has been found capable of recording a fetal ECG signal. The quality of the fetal ECG signal has been found to be the same as that recorded from a standard fetal scalp electrode. |
| Calimero Toco CE Report | 11-Oct-2004 | M2735 | 58868 | - The transducer with the `all flat` bottom did not meet the requirements for a reliable measurement of uterine contractions. It seems as if the flat bottom has problems finding tight contact to the abdominal wall and is easy to shift position. - The transducer with the 2 mm embossment showed satisfying sensitivity (93 to 102%, dependent on the gain) and did not seem to cause any discomfort in our patients. XXXX has adapted the design for the current generation of Toco transducers (M2734/M2735). |
| Intrauterine Pressure (IUP) |  |  |  | Comparisons of the recordings from both the Koala® and Intran® Plus catheters showed similar mean baseline uterine tone, peak pressures, contraction frequency and duration. |
| Clinical feasibility study: measurement of maternal heart rate by an infrared sensor built into the Tocography transducer in the context of cardiotocography | 20-Aug-2010 | M2734B | 59011 | On the basis of data records to date, the new CTG system with a built-in measurement of the maternal heart rate by the “abdominal pulse sensor” is able to reliably determine the maternal heart rate with a high level of quality. If the quantity is sufficient, it offers added safety for monitoring fetuses in utero with the same degree of comfort for the patient. The advantage of this additional measurement is clearly the added safety for the unborn child. |
| Toco MP M2734B Maternal Pulse Measurement | 17-Mar-2010 | M2734B | 59011 |  |
| M2702A/3A NBP Accuracy | 07-Feb-2007 | M2702/3A Avalon fetal monitors FM30/40 | 58869 | The obtained clinical data together with the laboratory validation data demonstrate that the non-invasive blood pressure (NBP) parameter of the XXXX Avalon FM20/FM30 Fetal / Maternal Monitor (M2702A, M2703A) meets the requirements (mean error of less than 5 mmHg and a standard deviation of less than 8 mmHg) over the specified blood pressure and pulse rate range. |
| Calimero CL – Clinical Evaluation Report | 15-May-2014 | Temperature | 83980 |  |
| Calimero CL – Clinical Evaluation Report | 15-May-2014 | Avalon Cableless Base Station (866074), Avalon Cableless Toco+ MP Transducer (866075), Avalon Cableless Ultrasound Transducer (866076), Avalon Cableless ECG/IUP Transducer (866077) | 83980 |  |
| Avalon CL Fetal & Maternal Patch and Pod | 11-May-2017 | Monica Patch | 138974 |  |

A study carried out by Monica Healthcare has demonstrated the equivalence of the MONICA PATCH fetal and maternal monitor (previous commercial name of Novii) in terms of accuracy and reliability compared to the FDA approved Monica AN24 predicate device:

| #### Document number | #### 107-CT-001 (Monica Healthcare) |
| --- | --- |
| #### Date of publication | #### 11-May-2017 |
| #### Product | #### Monica Patch |
| #### TFS# | #### 138974 |

Appraisal:

The clinical data collected from XXXX  initiated studies referenced in the Table above are fulfilling following suitability criteria (as of Appendix A, “Clinical Data Appraisal Methods”):

D1: data generated from actual device

A1: device used for the same intended use

P1: data generated from a patient group that is representative of the intended treatment population

R2: data contain sufficient information (with some deficiencies)

Note: Optimum Combination is: D1 + A1 + P1 + R1

And following contribution criteria:

T1: functionality in question directly mentioned by the study

O1: outcome measures reported reflect the intended performance / intended use

F2: no follow-up (not a therapeutic device)

S2: statistical significance (data include quantitative / qualitative statements, but statistical analysis is not always provided)

C1: the functionality in question used for day-to-day service and was clinically significant

Note: Optimum combination: T1 + O1 + F1 + S1 + C1

## Data reporting

All literature retrieved and selected will be retained and fully documented (including the selection and appraisal outcome) for incorporation in the Clinical Evaluation for this device.

# Appendix B – Plan for the Fulfillment of GSPR (General Safety and Performance Requirements)

| General Safety and Performance Requirements | Applicable Norms and Standards | Qualification - Elements of the Clinical Evaluation which will support the Fulfillment of this GSPR | Corresponding MDD ER |
| --- | --- | --- | --- |
| 1. Performance and safety Devices shall achieve the performance intended by their manufacturer and shall be designed and manufactured in such a way that, during normal conditions of use, they are suitable for their intended purpose. They shall be safe and effective and shall not compromise the clinical condition or the safety of patients, or the safety and health of users or, where applicable, other persons, provided that any risks which may be associated with their use constitute acceptable risks when weighed against the benefits to the patient and are compatible with a high level of protection of health and safety, taking into account the generally acknowledged state of the art. | EN ISO 14971:2017 Risk Management ISO 14971 ISO 13485 EU-MDR Annex 1 Chapter I: GSPR 1 | Intended purpose and intended use statement from IFU Clinical data from manufacturer and analysis of the clinical data Literature Reviews and PMS data appraisal | ER1 |
| 4. Reduction of Risks Risk control measures adopted by manufacturers for the design and manufacture of the devices shall conform to safety principles, taking account of the generally acknowledged state of the art. To reduce risks, Manufacturers shall manage risks so that the residual risk associated with each hazard as well as the overall residual risk is judged acceptable. In selecting the most appropriate solutions, manufacturers shall, in the following order of priority: a) eliminate or reduce risks as far as possible through safe design and manufacture; b) where appropriate, take adequate protection measures, including alarms if necessary, in relation to risks that cannot be eliminated; and c) provide information for safety (warnings/precautions/contra-indications) and, where appropriate, training to users. Manufacturers shall inform users of any residual risks. | ISO 13485 EN ISO 14971:2017 Risk Management: ISO 14971 EU-MDR Annex 1 Chapter I: GSPR 4 | State of the Art search Warnings and contraindications listed in the IFU Risk management analysis Review of labeling (IfU) | ER 2 |
| 8. Risk benefit ratio All known and foreseeable risks, and any undesirable side-effects, shall be minimized and be acceptable when weighed against the evaluated benefits to the patient and/or user arising from the achieved performance of the device during normal conditions of use. | ISO 13485 EN ISO 14971 :2017 EU-MDR Annex 1 Chapter I : GSPR 8 | Risk management analysis Benefits description Analysis of the acceptability of undesired side effects Generic risk profile Risk / Benefit assessment in CER | ER 6 |

Source: EU MDR (EU 2017/745) Annex I General Safety and Performance Requirements for inclusion in CE reports

Effective Date is the date of last signature.